NATIONAL INSTITUTE FOR HEALTH AND CARE EXCELLENCE

Medical technology consultation document

Rezum for treating lower urinary tract symptoms secondary to benign prostatic hyperplasia

The National Institute for Health and Care Excellence (NICE) is producing guidance on using Rezum for treating benign prostatic hyperplasia in the NHS in England. The medical technologies advisory committee has considered the evidence submitted by the company and the views of expert advisers. This document has been prepared for public consultation. It summarises the evidence and views that have been considered, and sets out the recommendations made by the committee. NICE invites comments from the public. This document should be read along with the evidence (see the committee papers). The advisory committee is interested in receiving comments on the following: Has the term ‘benign prostatic hyperplasia’ been used appropriately in the document to reflect the indication for the Rezum procedure? If you think another term should be used, please comment accordingly. Has all of the relevant evidence been taken into account? Are the summaries of clinical and resource savings reasonable interpretations of the evidence? Are the recommendations sound and a suitable basis for guidance to the NHS? Are there any equality issues that need special consideration and are not covered in the medical technology consultation document? Note that this document is not NICE's final guidance on Rezum for treating lower urinary tract symptoms secondary to benign prostatic hyperplasia. The recommendations in section 1 may change after consultation. After consultation the committee will meet again to consider the evidence, this document and comments from the public consultation. After considering the comments, the committee will prepare its final recommendations which will be the basis for NICE’s guidance on the use of the technology in the NHS in England. For further details, see the medical technologies evaluation programme process and methods guides. The key dates for this guidance topic are: Closing date for comments: 8 January 2020 Second committee meeting: 24 January 2020 Details of the advisory committee are given in section 5.

NICE medical technologies guidance addresses specific technologies notified to NICE by companies. The ‘case for adoption’ is based on the claimed advantages of introducing the specific technology compared with current management of the condition. This case is reviewed against the evidence submitted and expert advice. If the case for adopting the technology is supported, the specific recommendations are not intended to limit use of other relevant technologies that may offer similar advantages. If the technology is recommended for use in research, the recommendations are not intended to preclude the use of the technology in the NHS but to identify further evidence which, after evaluation, could support a recommendation for wider adoption.

Recommendations

Evidence supports the case for adopting Rezum for treating lower urinary tract symptoms (LUTS) caused by benign prostatic hyperplasia (BPH) in the NHS. Rezum relieves LUTS and improves quality of life.

Rezum is a minimally invasive procedure and can be done under local anaesthesia or light sedation. It should be considered as a treatment option for people with:

moderate to severe LUTS (International Prostate Symptoms Score [IPSS] typically 13 or over) and

a moderately enlarged prostate (typically between 30 cm3 and 80 cm3).

Cost modelling estimates that, per person over 4 years, Rezum is cost saving compared with:

UroLift by £497

transurethral resection of the prostate (TURP) by £569

holmium laser enucleation of the prostate (HoLEP) by £651.

Why the committee made these recommendations

Rezum is a minimally invasive procedure that involves injecting steam to destroy excess prostate tissue. Clinical evidence shows that using the Rezum procedure relieves LUTS caused by BPH in men with moderate to severe symptoms who have a moderately enlarged prostate. Evidence also shows that using Rezum is associated with improved quality of life and preserved sexual function. Cost analyses suggest that when Rezum is used as an alternative to standard treatment options such as TURP, it is likely to lead to cost savings because it is done as day surgery with reduced operating and recovery costs. Rezum is also cost saving compared with UroLift because of low consumable costs.

The technology

| Technology | Rezum is water vapour (steam) therapy for treating lower urinary tract symptoms (LUTS) associated with benign prostatic hyperplasia (BPH). The technology uses water vapour to destroy excess prostate tissue with the aim of relieving symptoms. The water vapour is injected through a single-use device attached to a urological endoscope into the prostate. The process is intended to disrupt cell membranes, leading to cell death and shrinking the prostate. The intention is to relieve obstructive symptoms without interfering with surrounding tissues that might impair sexual function. The vapour is injected for 9 seconds during treatment. The number of times this has to be done in each lobe of the prostate depends on the length of the prostatic urethra. It can be customised to the configuration of the gland. A maximum number of 15 full treatments can be done with each delivery device. The procedure is done under local, regional or general anaesthesia and lasts up to 20 minutes. |
| --- | --- |
| Innovative aspects | Rezum differs from other prostate treatments because it uses water vapour thermal energy. It does not use a laser and can be used to treat both the median lobe and an enlarged central zone. |
| Intended use | Rezum is intended for the treatment of prostates with volumes greater than 30 cm3 (equivalent to 30 g). The instructions for use state that Rezum is contraindicated for patients: • with a urinary sphincter implant • who have a penile prosthesis. |
| Costs | The typical consumable cost of the Rezum procedure is estimated at £1,348 (excluding VAT) per treatment. The company supplies the generator, which is loaned free of charge. The company also provides servicing including maintenance and other services (i.e. software update) free of charge. |
| For more details, see the website for Rezum. |  |

Evidence

## Clinical evidence

### Relevant evidence comes from 4 studies presented in 10 publications, including 1 randomised controlled trial

Four studies were relevant to the decision problem in the scope:

1 randomised controlled trial (5 publications: McVary et al. 2019, McVary and Roehrborn, 2018, Roehrborn et al. 2017, McVary et al. 2016a, McVary et al. 2016b)

1 prospective observational study (3 publications: Mynderse et al. 2015; Dixon et al. 2015, Dixon et al. 2016)

2 retrospective observational studies (Mollengarden et al. 2018, Darson et al. 2017).

- The randomised controlled trial was in 197 people with an International Prostate Symptom Score (IPSS) of 13 or more and an estimated prostate volume between 30 cm3 and 80 cm3, who had had no previous surgical interventions for their prostate. The observational studies included people with prostate sizes from 20 cm3 to 110 cm3 who underwent the Rezum procedure. All are non-UK studies.

### The evidence shows that Rezum is clinically effective

The Rezum II study showed that Rezum was associated with statistically significant improvements in lower urinary tract symptoms (LUTS) compared with sham at 3-month follow up. These improvements were sustained throughout 4 years of follow up. The treatment benefits of Rezum in relieving LUTS were also consistently observed in the observational studies. These improvements were gained without significantly adversely affecting sexual function, at least in the short term. Overall, the evidence base shows that Rezum is an effective treatment for LUTS in people with benign prostatic hyperplasia (BPH). Rezum also improved measures of quality of life (McVary et al. 2019, Darson et al. 2017; Dixon et al. 2015 and 2016).

### There is no evidence that directly compares Rezum with other interventions for BPH

None of the included studies compared Rezum with other commonly used treatments for BPH. Clinical experts suggested that more invasive treatments such as transurethral resection of the prostate (TURP) were likely to be associated with a more substantial relief of urinary symptoms than Rezum. But there is currently no direct evidence to support this. Similarly, there are no direct comparisons of Rezum with UroLift, holmium laser enucleation of the prostate (HoLEP), or GreenLight laser. Expert opinion indicated that recruiting participants to clinical trials that directly compare different minimally invasive and invasive treatments is challenging because people often express a preference to avoid more invasive treatment.

### An indirect comparison suggests that Rezum is as effective as UroLift

In the absence of direct comparative evidence, the company did an indirect comparison of Rezum and UroLift to relieve LUTS. This was based on the results of the Rezum II study and the Luminal Improvement Following prostatic Tissue (LIFT) study (Roehrborn et al. 2017b). Both technologies are minimally invasive procedures to treat LUTS, and the trial designs and study populations were similar. The main exception was that the Rezum II study included people with median lobe obstruction (31.1% of study participants) while the LIFT study did not. A comparison of the results from the 2 trials indicated that the therapeutic effects of Rezum and UroLift in relieving LUTS were similar, but people were less likely to need further operations after Rezum than after UroLift. The retreatment rate was 4.4% at year 4 with Rezum and 13.6% at year 5 with UroLift.

### Rezum is considered a safe procedure

The Rezum II study reported 3 procedure-related serious adverse events in the 3-month follow up, including extended urinary retention, and nausea and vomiting, which were considered to be due to the sedative medication. An additional 3 procedure-related serious adverse events were reported with Rezum during the 3 to 12-month follow-up period, including bladder contracture, bladder stone, and urosepsis after cystoscopy. The clinical experts did not identify any specific safety concerns with Rezum.

## Cost evidence

### The company suggests that using Rezum is cost saving compared with other treatments for BPH

The company developed a decision analytic model with a time horizon of 4 years. The model compared Rezum with 4 comparators: UroLift, GreenLight laser, HoLEP and TURP. The model assumed that all the technologies had equal efficacy in alleviating LUTS associated with BPH. The model incorporated a cohort Markov structure. Erectile dysfunction and urinary incontinence were included as permanent adverse events that inform long-term health states. The need for surgical retreatment for recurrence of LUTS was also considered. The results of the company model indicated that Rezum was cost saving by £737, £532, £25 and £758 per person when compared to TURP, UroLift, GreenLight and HoLEP respectively over 4 years.

### The EAC’s changes to the cost model more accurately reflect empirical evidence and expert opinion

The main parameters in the model were the technology costs, theatre time, hospital length of stay, adverse events and the need for another operation. The external assessment centre (EAC) adjusted some of the model’s parameters, including the surgical retreatment rate and the adverse event rates, to more accurately reflect published empirical data and expert opinion.

### Rezum is cost saving compared with TURP, UroLift and HoLEP but cost neutral compared with GreenLight

The EAC base case results showed that Rezum was cost saving by £569, £497 and £651 compared with TURP, UroLift, and HoLEP per patient respectively over 4 years. Rezum remains cost saving when all parameters are subjected to a one-way deterministic sensitivity analysis. In the base case, Rezum is cost incurring by £62 per patient over 4 years compared with GreenLight laser. The model assumed that GreenLight, like Rezum, was used as a day case. If, in practice, this is not the case, then Rezum is anticipated to be cost saving. Overall, the EAC considered Rezum, therefore, to be approximately cost neutral compared with GreenLight over the course of 4 years.

Committee discussion

## Clinical-effectiveness overview

### Rezum is an effective minimally invasive procedure with clinical benefits

The committee concluded that the evidence from the Rezum II study demonstrated the effectiveness of Rezum in relieving lower urinary tract symptoms (LUTS) associated with benign prostatic hyperplasia (BPH) with a sustained benefit up to 4 years of follow up. The committee noted that this is supported by the results of the observational studies and by the results of a NICE patient experience survey that was provided to the committee. The committee noted that there are no studies that directly compare Rezum with other treatments in relieving symptoms in people with BPH, but considered an indirect comparison between Rezum and UroLift, that was drawn from analogous trial data. This suggests that Rezum is at least as effective as UroLift over 4 years. The clinical experts explained that these 2 minimally invasive procedures are used in similar cohorts of populations in clinical practice and that, in their experience, both procedures provide a similar degree of symptom relief. They pointed out, however, that Rezum is more versatile than UroLift in treating different shapes of prostate, for example in men with an obstructive median lobe.

### Rezum should be used for men with moderate to severe LUTS with an estimated prostate volume of 30 cm3 to 80 cm3

The committee noted that there is one pivotal study that provides the evidence for the efficacy of Rezum. The clinical experts explained that Rezum II was a US study and designed to meet US Food and Drug Administration eligibility criteria. Its major inclusion criteria were: men aged at least 50 or over who have symptomatic BPH with an International Prostate Symptoms Score (IPSS) of 13 or greater, and with a prostate volume, measured by transrectal ultrasound, of 30 cm3 to 80 cm3. The committee concluded that there is limited evidence on the efficacy of Rezum in men outside this cohort. The clinical experts confirmed that, in their clinical practice, this cohort of patients corresponds closely to those that they treat with Rezum and that this encompasses approximately 75% to 85% of the overall population that need treatment to relieve LUTS. The clinical experts also explained that, for people with mild LUTS (IPSS less than 8), first-line treatment is medication or lifestyle change. For people with an estimated prostate volume 120 cm3 and greater, more invasive surgical interventions are recommended.

### Rezum is unlikely to damage surrounding tissue and nerves, and sexual function is not affected in the short term

The clinical experts explained that loss of sexual function is an important concern for people undergoing invasive treatment for LUTS because the invasive procedure is likely to cause damage to nerves on the external surface of the prostate. They also explained that Rezum involves injecting steam into carefully directed and localised areas of the prostate from the inner, urethral surface of the prostate, and this may avoid nerve damage. The committee considered that the published evidence suggests that sexual function is retained after treatment with the Rezum procedure. It did note, however, a high incidence of sexual inactivity in people included in the Rezum II study and that overall sexual function showed a tendency to decline during study follow up. This may be attributable to the effect of ageing in the study population. Overall, the committee concluded that sexual function is preserved with Rezum, at least in the short term, and that this may be particularly important to people who are sexually active at the time of treatment. The committee was uncertain, however, about the impact of Rezum on longer-term sexual function because no data are available for longer than 4 years.

### Quality of life is an important outcome when considering patient benefit

The evidence from the Rezum II study and observational studies indicated that treatment with Rezum with significant relief of LUTS is associated with a significant improvement in quality of life, which persists for up to 4 years of follow up. The clinical experts confirmed that, in their experience, people who underwent Rezum express a high level of satisfaction after the procedure.

## Side effects and adverse events

### Urinary tract infection is a common complication after Rezum

The clinical experts advised that complications that are encountered after the Rezum procedure include urinary tract infections (UTIs), bleeding, epididymitis and abscess. The clinical experts also explained that, after the Rezum procedure, a urinary catheter is left in place to allow the dead prostate tissue to drain away. The need for catheterisation, combined with the presence of necrotic tissue, are considered by the clinical experts to be predisposing factors for developing UTIs. This is higher for Rezum than UroLift, which does not need a post-operative urinary catheter. The clinical experts estimated that the risk of UTIs is around 5% to 7%, so a 5 to 7-day course of prophylactic antibiotics is usually prescribed after the procedure. The committee concluded that post-procedure UTI rates associated with Rezum may be difficult to record because patients may present to their GP for treatment. It also noted that antibiotic use was not reported in the Rezum II study.

### The rate of surgical reintervention is low with Rezum

The committee noted that the Rezum II study reported a 4.4% rate of surgical retreatment over 4 years of follow up. The clinical experts suggested that the average retreatment rate in their experience is around 3% after Rezum, and that retreatment is most likely in the first year after the procedure. The clinical experts explained that, because there is no direct view of the prostate cavity during the Rezum procedure, additional transurethral resection of the prostate (TURP) was sometimes needed to remove residual prostate tissue after Rezum. Overall, the committee noted that the retreatment rate with Rezum is low and compares favourably with similar treatments like UroLift.

## Relevance to the NHS

### The evidence for Rezum is broadly generalisable to the NHS

The clinical experts explained that Rezum is currently done in some NHS trusts and that there has been a popular demand by people for this procedure in some centres. The committee noted that the published evidence for Rezum is from studies that were done outside the UK. Nonetheless, the clinical experts explained that the study population included in the Rezum II study is similar to the people that they treat with Rezum in their own practice in the NHS. The committee concluded that the evidence is generalisable to UK NHS practice.

## NHS considerations overview

### Rezum is a day surgery procedure that can be done under local anaesthetic or light sedation but it may not be suitable for everyone

The clinical experts said there are currently 8 different treatments, including Rezum, available in the NHS for people with significant LUTS that have not responded to conservative therapy including medication and lifestyle changes. The clinical experts considered TURP to be the standard of care for LUTS secondary to BPH, but emphasised that treatments need to be offered to people on an individual basis guided by their individual circumstances. Key factors for consideration include: the availability of procedures in their local hospitals, age, prostate gland size and characteristics, and comorbidities. Rezum’s advantages over other technologies are that it is a minimally invasive procedure that can be done under local anaesthesia or light sedation and that takes only around 20 minutes. People are therefore usually offered day case treatment. The clinical experts said that Rezum should be avoided in people with prostatitis or confirmed prostate cancer, in people for whom day case treatment is impractical or unsafe, and if there’s a risk of increased bleeding, for example if they’re having anticoagulant treatment.

### Rezum is used to treat patients with benign prostate enlargement but there is no consensus on how to measure prostate size

The clinical experts said that an enlarged prostate that causes LUTS as a result of prostatic obstruction is caused by prostatic hyperplasia, which is a benign histopathological diagnosis. The clinical experts explained that there is currently no consensus on how prostate size should be estimated or measured in UK clinical practice. They considered that normally imaging would be used to estimate prostate size before surgically invasive treatment. The clinical experts said that imaging modalities could complement information from rectal digital examination of the prostate. Common imaging tools include transrectal ultrasound, cystoscopy and MRI. On the basis of these measurements, the committee heard that Rezum is usually offered to people with moderate prostatic enlargement with a prostate that is typically estimated to be 30 cm3 to 80 cm3.

### The Rezum procedure is easy to learn

The clinical experts explained that urologists need specialist training to do the Rezum procedure. This training is provided by the company and includes lectures and simulation training. The clinical experts suggested that Rezum is relatively easy to learn and that the training requirement is minimal.

## Cost modelling overview

### Rezum is cost saving compared with other treatments for BPH but there are limitations in the cost model

The committee noted that the external assessment centre’s (EAC) cost modelling results showed that Rezum is likely to be cost saving compared with TURP, UroLift, and holmium laser enucleation of the prostate (HoLEP) by more than £497 per patient over 4 years. The committee noted, however, that there are some limitations in the model, including the assumption that all treatments are equally effective in relieving LUTS. Indirect comparative data from the trials suggest that the technologies may not all reduce the IPSS score to the same extent. The clinical experts confirmed that more invasive procedures such as TURP, which removes prostate tissue, would be expected to have greater IPSS improvements. It’s uncertain to what extent this impacts the need for retreatment. The committee identified other limitations in the cost model, including the fact that no consideration was given to the impact of urinary catheterisation and removal, or the need for antibiotics after Rezum. Nonetheless, the committee considered that these were unlikely to substantially affect Rezum’s cost savings.

## Main cost drivers

### Doing Rezum as day surgery is the main driver for cost savings

Rezum is commonly done as day surgery and people are not usually admitted to hospital afterwards. The EAC considered that this was a key driver in the estimated cost savings when Rezum compared with traditional treatment options such as TURP. The company’s model showed that the consumable cost of Rezum was estimated to be around £1,348 per person. The company provides the generator and servicing such as maintenance free of charge This cost relative to competitor treatments was also influential in the cost modelling results. The company representatives confirmed that they do not anticipate any changes to this cost model for the foreseeable future.

## Cost savings

### Rezum is cost saving compared with other treatments for BPH

The EAC did deterministic sensitivity and probability sensitivity analyses that varied parameters in cost models, and the results showed that Rezum remained cost saving. The committee concluded that, based on the published evidence, cost modelling and expert opinion, using Rezum is likely to lead to a cost of £497 compared with UroLift, £570 compared with TURP, and £650 compared with HoLEP, for every patient treated over a 4-year time horizon.

## Further research

- Further research will improve the evidence base for the efficacy of Rezum compared with surgical treatment

Further evidence to address the efficacy of Rezum compared with other surgical interventions would be welcome, including their relative impact on symptom relief, quality of life and preservation of long-term effect sexual function.

Committee members and NICE project team

## Committee members

This topic was considered by the medical technology advisory committee which is a standing advisory committee of NICE.

Committee members are asked to declare any interests in the technology to be appraised. If it is considered there is a conflict of interest, the member is excluded from participating further in that evaluation.

The minutes of each committee meeting, which include the names of the members who attended and their declarations of interests, are posted on the NICE website.

## NICE project team

Each medical technologies guidance topic is assigned to a team consisting of 1 or more technical analysts (who act as technical leads for the topic), a technical adviser and a project manager.

### Ying-Ying Wang
Health Technology Assessment Analyst

### Paul Dimmock and Bernice Dillon
Health Technology Assessment Advisers

### Elizabeth Islam
Project manager

### ISBN: [to be added at publication]