NATIONAL INSTITUTE FOR HEALTH AND CARE EXCELLENCE

HEALTHTECH PROGRAMME

Draft guidance

Intermittent urethral catheters for chronic incomplete bladder emptying in adults: late-stage assessment

Guidance development process

NICE late-stage assessment (LSA) guidance evaluates categories of technologies that are already in widespread use within the NHS. It assesses whether price variations between technologies in a category are justified by differences in innovation, clinical effectiveness and patient benefits. This will support NHS commissioners, procurement teams, patients and healthcare professionals to choose technologies that maximise clinical effectiveness and value for money.

Find out more on the NICE webpage on late-stage assessment (LSA) for medtech.

NICE is producing this guidance on intermittent urethral catheters for long-term urinary management in the NHS in England. The medical technologies advisory committee has considered the evidence and the views of clinical and patient experts.

This document has been prepared for consultation with the stakeholders. It summarises the evidence and views that have been considered and sets out the recommendations made by the committee. NICE invites comments from the stakeholders for this evaluation and the public. This document should be read along with the evidence.

The committee is interested in receiving comments on the following:

Has all of the relevant evidence been taken into account?

Are the summaries of clinical and cost effectiveness reasonable interpretations of the evidence?

Are the recommendations sound and a suitable basis for guidance to the NHS?

Are there any aspects of the recommendations that need particular consideration to ensure we avoid unlawful discrimination against any group of people on the grounds of age, disability, gender reassignment, pregnancy and maternity, race, religion or belief, sex or sexual orientation?

After consultation:

Based on the consultation comments received, the committee may meet again.

If committee meets again it will consider the evidence, this evaluation consultation document and comments from stakeholders.

The committee will then prepare the final draft guidance, which will go through a resolution process before the final guidance is agreed.

Note that this document is not NICE's final guidance on Intermittent urethral catheters for chronic incomplete bladder emptying in adults. The recommendations in section 1 may change after consultation.

More details are available in NICE’s health technology evaluations: the manual and NICE’s late-stage assessment interim process and methods statement.

Key dates:

Closing date for comments: 7 March 2025

Second committee meeting: 17 April 2025

Recommendations

  - There is not enough evidence to determine whether price variations between different intermittent urethral catheters are justified. So, if more than one catheter is available use the least expensive catheter that:

is clinically appropriate, and

meets the preferences and needs of the person with chronic incomplete bladder emptying.

  - A healthcare professional and the person with chronic incomplete bladder emptying should decide together which intermittent urethral catheter to use. Decisions should take into account:

ease and comfort of insertion

risk of infection.

  - Urology services should have access to a range of intermittent urethral catheters available for prescription in the NHS, so that adults with chronic incomplete bladder emptying can have the most appropriate catheter for them.

## What information is needed

More information is needed to show if price variation between different intermittent urethral catheters can be justified and attributed to any specific feature.

Evidence should compare catheter features with each other to show if a specific feature affects outcomes and user preferences. Evidence should be generated across different groups of people who use intermittent catheters for bladder drainage, through formal research studies or real-world evidence.

A core outcome set, including validated patient-reported outcomes, should be developed so that outcomes can be reported consistently. Features should be classified consistently so they can be assessed in a standardised way.

Key outcomes and information that should be captured include:

details of the catheter features

incidence of urinary tract infection

incidence of haematuria, in particular macroscopic or visible haematuria

residual urine volume (the volume of urine remaining in the bladder after catheterisation)

comfort during catheterisation (insertion and removal)

ease of use

health-related quality of life

adverse events

how a feature contributes to an outcome.

## What this means in practice Considerations for procurement and commissioning In 2023 to 2024, a total of 95,437,405 intermittent catheters were prescribed in the NHS, and the price per catheter ranged from £0.40 to £3.28. Many factors can influence which type of intermittent catheter is most appropriate to use and how effective it might be. Some features might be needed for specific groups or people in specific situations. For example, people with mobility issues might benefit from catheters with an integrated drainage bag. People with reduced manual dexterity might need catheters with specially designed packaging for ease of opening or catheters with handles or grippers for ease of handling. Features that meet clinical needs and user preferences may be worth paying more for. But not all features are associated with higher prices. If a company introduces a catheter with new features to the market, evidence of clinical superiority should be provided to show additional value. Evidence should also show how it makes a difference in outcomes and contributes to people’s needs and preferences. Considerations for healthcare professionals Decide together with the person with chronic incomplete bladder emptying which catheter to use, following the principles of NICE’s guidance on shared decision making. Offer a choice of catheters to the person with chronic incomplete bladder emptying, and offer training on how to use each type of catheter. It may be appropriate to prescribe more than one type of catheter to be able to meet the person’s needs and lifestyle in different settings and situations. Information for people with chronic incomplete bladder emptying People with chronic incomplete bladder emptying should be given a choice of catheter and training on how to perform intermittent catheterisation. If the person is using a catheter that causes complications, such as pain or discomfort, urinary tract infection or bleeding, they should be offered support to see if changing the catheter type will help.

## Why the committee made these recommendations

Intermittent catheters are used for catheterisation for medical bladder emptying, particularly for long-term bladder management. There are many intermittent catheters available, which vary in features and cost. Most variation in prices of intermittent urethral catheters can be accounted for by basic and additional features. This assessment aimed to determine whether the differences in clinical, economic and non-clinical outcomes attributed to those features could justify price variation.

There is no robust clinical evidence of the effectiveness of individual features of catheters. The available evidence does not consistently report on the most important outcomes. Where there is limited evidence for a particular feature, there is a lack of information about how the reported outcome can be attributed to that feature. User-centred work shows that the most important factors for people who use intermittent catheters for bladder drainage are ease and comfort of insertion and risk of infection. It also shows that people’s needs, preferences and experiences of using intermittent catheters varied.

There is not enough evidence to determine whether price variations between catheters with different features are justified, and more information is needed.

The technology

The technology identified for this assessment is single-use and sterile intermittent urethral catheters for bladder drainage. Intermittent catheters are used for catheterisation for medical bladder emptying, particularly for long-term (more than 28 days) bladder management. Intermittent catheterisation is usually done several times a day.

The NHS Supply Chain bladder and bowel framework describes an intermittent catheter as ‘a smooth, flexible tube with holes, used for short term drainage of urine from the bladder and the catheter has no balloon’. Intermittent catheters can be made of different materials, are available in different sizes, and have different coatings and tips. The basic requirements of single-use and sterile catheters are detailed in the scope.

This feature-based assessment included a total of 838 intermittent catheters, across 86 catheter product lines (aggregated by catheter size, male or female, tip type) from 17 manufacturers. Based on the NHS Drug Tariff (part IXA), the price of each catheter ranged from £0.40 to £3.28. This indicated that the maximum price difference could be £2.88 per catheter.

To determine whether pricing variation is justified, this assessment considered 8 additional features (detailed in the scope):

integrated drainage bag

integrated handle or markings

tip protector or introducer

microhole zone technology

enhanced lubrication or coating (referred to as ‘enhanced coating’ in the assessment)

specially designed catheter case

specially designed packaging.

- Committee discussion

The medical technologies advisory committee considered evidence on additional features of intermittent urethral catheters for chronic incomplete bladder emptying in adults from several sources. The evidence included clinical evidence from targeted literature searches, evidence from company submissions and expert feedback. The committee also considered the economic evidence from a review of the published literature and the evidence submitted by the companies, an economic evaluation done by the external assessment group (EAG), and 3 pieces of user-centred work. Full details are available in the project documents for this guidance.

## The condition

  - Chronic incomplete bladder emptying can be caused by neurological or non-neurological conditions. People with this condition need to use intermittent catheters for long-term bladder management.

## Current practice

NICE’s guideline on the management of lower urinary tract symptoms in men recommends that intermittent catheterisation should be offered for bladder drainage before indwelling urethral or suprapubic catheterisation.

Intermittent catheterisation can be done by adults of any age. The Royal College of Nursing’s guidance on catheter care recommends that people who perform catheterisation should be educated about and trained in this procedure, and all training should be done by a healthcare professional who is competent in providing training. The European Association of Urology Nurses’ guideline on urethral intermittent catheterisation recommends follow-up training and ongoing support. The frequency of performing intermittent catheterisation varies depending on the individual’s presenting symptoms. The frequency should be reviewed regularly, as symptoms can improve or deteriorate depending on individual circumstances. But, there can be variation in the quality of training, support offered and review provided.

## User experiences, needs and preferences

Patient experts talked about their experiences of using intermittent catheters for bladder drainage and the challenges they faced. They felt that doing intermittent catheterisation is a lifestyle change. They also explained their needs and highlighted the importance of choice and informed decision making. Patient experts emphasised the need to adhere to catheterisation to prevent infection, and the importance of comfort during catheterisation from insertion to removal.

User-centred work by NICE included a thematic review of the literature, a user preference assessment and an online survey. This user-centred work suggested that people’s needs, preferences and experiences of using intermittent catheters varied. The most common factors that people considered important when selecting a catheter were ease and comfort of insertion, and risk of infection. Different catheters need to be available to meet people’s lifestyles and needs across a range of situations and settings. People faced various challenges, and they were more likely to experience difficulties when using a catheter away from their home. Shared decision making was considered important when choosing a catheter, but in practice choice was not always available.

The committee discussed the evidence from the user-centred work. It acknowledged the various needs of individuals in different settings and situations, and highlighted the importance of shared decision making. The clinical and patient experts agreed it was important to empower people and offer them a choice, and for clinicians to be aware of the range of catheters available. The clinical experts highlighted that regular review of catheter suitability is key to ensuring that the most appropriate catheter is used to meet people’s needs and preferences. But, the committee noted that the user preference assessment lacked data on the preferences of people with mobility issues.

## Equality considerations

The needs of people with chronic incomplete bladder emptying vary from person to person. Access to a wide range of intermittent catheters is needed to meet different people’s needs. Additional support or adaptations may be needed to enable people who would otherwise not benefit from intermittent catheterisation to use this procedure. For example, people with mobility issues might benefit from using catheters that have an integrated bag or can be connected to an external collection bag. The committee noted that there was no evidence of the clinical benefit of an integrated bag feature, but it agreed that this feature might be needed for specific groups or people in specific situations.

## Clinical effectiveness

### Key evidence and feature categories

The clinical review included 18 comparative studies evaluating intermittent catheters with the additional features. Across the 18 studies, from 1997 to 2024, there were 3 randomised controlled trials (RCTs), 12 crossover RCTs, 1 cross-sectional study, 1 prospective comparative cohort study, and 1 in vitro study. The follow-up durations varied across studies from 1 day to 1 year. There was variation in who performed the intermittent catheterisation across studies. Most studies were company funded or included authors who were company employees.

The included evidence covered all additional features identified. There was a difference between the feature of focus (hypothesis tested) of the study and the features present within that study. Most studies aimed to compare coated and uncoated catheters. The features of the interventions and comparators varied across studies, and some studies compared more than 2 types of catheter. There was substantial variation in outcome definitions, measures and reporting.

The committee discussed the variations in features and outcomes reported across studies, and understood that there was a lack of consensus about how features were categorised. The committee heard that there is currently no agreement on categorisation of features for intermittent catheters.

The committee recognised that some of the evidence was outdated, and this might not reflect current practice.

The committee acknowledged the lack of evidence on the effectiveness of individual features. It agreed that a component network meta-analysis (CNMA) was an appropriate method to analyse the benefits of the components.

### Results of the CNMA

The clinical evidence reported on 5 key outcomes, but the CNMA was feasible for 3 outcomes and isolated the effects of 4 features. When compared with a basic uncoated catheter, there was a possible reduction in urinary tract infection (UTI) with enhanced coating or specially designed packaging. Microhole zone technology may reduce the risk of haematuria (mainly relating to microscopic haematuria), whereas specially designed catheter cases may increase the risk of haematuria. But the committee noted there was no information about how a specially designed catheter case could have this effect and questioned the validity of this result. For the other 4 features (integrated drainage bag, integrated handle or marking, insertion sleeve or grip, tip protector or introducer), there was insufficient evidence to isolate their effectiveness.

The committee noted that the number of studies included in the analysis was limited because of the variation in features and outcomes reported across studies. It also highlighted the potential benefits of microhole zone technology and specially designed packaging, but the certainties were ‘very low’ across all CNMA outcomes. The committee discussed the clinical importance of microscopic and macroscopic haematuria. It agreed that only macroscopic and visible haematuria should be considered clinically important.

### Uncertainty about the effectiveness of individual features

The committee acknowledged the lack of consensus on the feature categories, heterogeneity of the data, uncertainties of the outcomes, and that some evidence was outdated. It concluded that, although there is evidence for catheters with different features, there is not enough evidence of the effectiveness of individual features separately, or how each feature contributes to the reported outcomes. So, more information is needed to estimate the effect size of individual catheter features. This should include the details of the features, and plausible relationships between the features and outcomes.

## Cost effectiveness

### Regression analysis

Catheters listed on the NHS Drug Tariff have multiple additional features and basic features. To estimate the average cost of individual features, a multiple regression model was used based on the price of 838 intermittent catheters. The model suggested that approximately 78% of the variation in prices could be accounted for by the presence of basic and additional features. But, based on one-way sensitivity analysis, the results of the regression model were sensitive to changes. The committee noted that not all features contributed to price variation and led to price increases. But, it understood that the cost of each feature that contributed to the price of a catheter might not be a true representation, because market intervention might change feature prices.

### Economic model structure

The EAG developed a multi-state Markov model to estimate the cost-effectiveness of individual catheter features. This was informed by an economic model used in NICE’s guideline on prevention and control of healthcare-associated infections. The model used a time horizon of 2 years and was restricted to people with urinary retention. The committee considered the structure of the model to be appropriate.

### Model assumptions

The model assumed that people used the same type of catheter over the 2-year time horizon. The committee noted that this assumption did not reflect people’s experience. The survey results showed that most people used intermittent catheters for more than 2 years and used more than one type of catheter. The committee understood that the model focused on features alone rather than real-world analysis, and there was very limited quality-of-life data and no long-term evidence. Sensitivity analysis showed that longer time horizons (up to a lifetime) had little impact on the direction or magnitude of effect, because most of the costs attributed to the catheter and differences in quality-adjusted life-years (QALYs) were small.

The committee discussed the use of assumptions for the quality-of-life and ease-of-use parameters. The EAG explained that there were different ways of measuring ease of use, and it was challenging to attribute it to individual features. It added that there was very limited data for quality of life, and that ease of use impacted on quality of life indirectly. The committee recognised the uncertainty in these assumptions and noted it as a limitation of the model.

The committee queried the standardised mortality ratio of 5.41 used in the model, because it was used for people recovering from spinal cord injury. But, it recognised that mortality was not a key driver for the analysis because of the model’s relatively short time horizon of 2 years.

### Results of the economic model

Over a 2-year time horizon, the base-case model found that 4 features (integrated handle or markings, insertion sleeve or grip, microhole zone technology, and specially designed packaging) were dominant (less costly and increased QALYs). The model also found that 1 feature (specially designed catheter case) was less costly and slightly less effective, but this feature was considered cost effective at the £20,000 per QALY threshold. The cost effectiveness of enhanced coating was uncertain, and integrated drainage bag and tip protector were not considered cost effective at a £20,000 willingness-to-pay threshold. Extensive one-way sensitivity analysis did not markedly change these findings.

The committee considered the model’s results and noted that all modelling was restricted to people with urinary retention. Given a lack of robust evidence and the assumptions used, the model’s results were subject to parameter uncertainty. The committee concluded that the model’s results needed cautious interpretation, but acknowledged that not all features correlated with price variation.

### Resource impact

The committee considered a resource impact assessment (RIA) that calculated the financial impact of catheter features. The RIA used the results of the EAG’s multiple regression model. It showed that the integrated bag feature had the highest annual cost per person, followed by tip protector and enhanced coating. Three features (insertion sleeve or grip, micro-hole zone and specially designed catheter case) had a zero cost. The committee recalled that the cost of each feature from the regression model might not be a true representation, so the costs or savings from the scenario analysis might not be realisable.

## Justification for price variation

The committee discussed the clinical and economic evidence and the work on user preference. It concluded that it was not possible to determine whether the differences in costs between intermittent catheters were justified by benefits derived from additional features based on the evidence. It emphasised the importance of shared decision making and offering a choice of catheter, as well as having access to a range of catheters with different features, so that people with chronic incomplete bladder emptying can have the appropriate catheters in different settings and situations.

## Evidence needed to show additional value

The committee concluded that more evidence is needed to justify the price variation between intermittent catheters with additional features. It acknowledged the limited evidence for the effect of individual features, and that it was unclear how particular features improved clinical outcomes and related to user needs and preferences. So, the committee noted that further evidence should particularly include details of the features, and plausible relationships between features, reported outcomes and user preferences.

Committee members

This topic was considered by NICE's medical technologies advisory committee, which is a standing advisory committee of NICE.

Committee members are asked to declare any interests in the technology to be evaluated. If it is considered there is a conflict of interest, the member is excluded from participating further in that evaluation.

The minutes of each committee meeting, which include the names of the members who attended and their declarations of interests, are posted on the NICE website.

### Specialist committee members

Faizan Awan
Patient specialist committee member

Daphne Colpman
CNS – continence, tertiary urology, University College London Hospitals NHS Foundation Trust

Marcus Drake
Professor of neurological urology, Imperial College London

Mandy Fader
Professor of continence technology, University of Southampton

Sarah Markham
Patient specialist committee member

Laura Rose
Clinical team leader, Continence Service NHS Ayrshire and Arran

### Clinical and patient experts

Soosie Cole
Clinical lead benign urology nurse specialist, Royal Devon University Hospital Foundation Trust

Sam Davison
Clinical procurement lead nurse, South Tees Hospitals

Austen El-Osta
Director, Self-Care Academic Research Unit (SCARU)

Jacqueline Emkes
Patient expert

Keith Ford
Patient expert

Robin Ward
Publicly elected governor (and lay member for Manchester Primary Care), NCA Foundation Trust & Greater Manchester Primary Care Digital Board

Ann Yates
Director of continence service – nursing, Cardiff and Vale University Health Board

ISBN: [to be added at publication]